Customer Service Department Fast Fax

To:   User		From:  Customer Service Department

2 Pages total including this page.  Contact us immediately if pages are missing or not legible.

Proper Mounting of the FARO Arm

Mounting the Base

The counterbalance by the tension spring generates torque at the base of the FaroArm so to achieve optimum machine accuracy, the mounting must meet certain requirements.

NOTE: The FaroArm must be mounted in a upright position. Do not mount the FaroArm in an inverted (upside down) or sideways position.

To fasten the FaroArm:

- Attach the 3.5" threaded ring and surface mount plate to any stable location. Tighten all mounting bolts to 100-inch pounds.

- Place the FaroArm on top of the 3.5” threaded ring.

- Screw the threaded collar clamp onto the base of the FaroArm and the 3.5” threaded ring.

- Use the wrench to tighten the threaded collar clamp.

Test the stability of the mount with a dial indicator. Attach a rigid bar, (approximately 250 millimeters/10 inches) to base of the FaroArm.  Remove either screw that holds the outer plastic cover in place, from the side of the FaroArm using a #4M hex wrench.

NOTE: The bar must have a 6M thread on one end to connect to the FaroArm. Measure the deflection of the bar when the FaroArm is fully extended. The angular deflection and lateral deflection of the base must not exceed the values in Table 1-1.

If you are still experiencing difficulties, please contact the customer service department by telephone at 1-800-736-2771 or by e-mail at support@faro.com.

FAX questions to:  (407) 333-8056