Customer Service Department Fast Fax

To:	FaroArm User			From:	Customer Service Department

Form:	176-015

2   Pages total including this page.  Contact us immediately if pages are missing or not legible.

TOPIC: FaroArm Loss of a Degree of Freedom

In the working volume of the FaroArm, there can be a loss of a degree of freedom (the natural rotation of the transfer case).  With this loss, bending will occur on the transfer tubes of the FaroArm causing a movement of the probe position which can not be recorded by the arm’s encoder system.  Measurement results taken in these positions will not be accurate.  This condition is not encountered in the calibration of the FaroArm.

Common areas of measurement in which this loss will occur are:

- When an encoder has reached a hard end-stop.

- Directly above the FaroArm.

- In close to the base of the FaroArm.

The FaroArm should always feel “fluid” in its movement.  If excessive force is needed to move to a measuring location, a degree of freedom has probably been lost.  The following figures illustrate some of the possible positions of a FaroArm where a degree of freedom has been lost.

Straight Up All FaroArm Models

Full Extension All FaroArm Models

90 Degrees, Up or Down 2-1-3 and 2-2-3 (with axis locked)

Inverted Arms – (Note:  In regards to the loss of a degree of freedom, the inverted arms are subject to the same restrictions as the normally mounted arms.)

As a general rule, if the Arm feels as though it is “locked-up”, or its motion does not feel fluid, you are probably experiencing a loss of a degree of freedom.

FAX questions to:  (407) 333-8056