Customer Service Department Fast Fax

To: FARO Arm User		From:  Customer Service Department

2 Pages total including this page.  Contact us immediately if pages are missing or not legible.

Proper use of the Vacuum Mount

This vacuum mount is meant for every size and model of USB Faro Arms and Faro Laser Tracker. Its purpose is to quickly, easily and rigidly mount the Faro Arm and Tracker to granite surfaces without degradation in the accuracy of the readings.

1. Vacuum Pump

2. Pressure Relief Valve

3. Flexible Tubing

4. 3 ½” Mounting Ring

5. Vacuum Mount

6. Pressure Gage

For proper use of the vacuum mount:

- Smooth clean granite surface, i.e. no dust, dirt or loose particles.

- Thin layer of the provided oil needs to be on the dirt free surface. The oil needs to be spread around evenly and as thin as possible, if there is too much oil it will slide.

- Make sure the Pressure Relief Valve on the pump is tightly screwed. If it is not screwed all the way, tighten it to allow for there to be a build up of pressure.

- Check seal for any cracks or splits.

- The initial pressure should be 30in Hg (which is in the blue area of the pressure).

- Check pressure gage periodically. Re-pressurize to 25 in Hg if the vacuum falls below 15 in Hg within 4 hours. On the gage there is color coding, green is the normal range between 15 and 25 in Hg, the red area is below normal, between 0 and 15in Hg.  If vacuum fails, please repeat the above steps, ensuring the mount is sitting on a clean flat surface, oil has properly been applied and that the seal is in good condition.

- FARO does not recommend leaving equipment mounted overnight with the Vacuum Mount.

For questions and concerns please email FARO Customer Service at:

North America       	Europe	Asia Pacific

250 Technology Park 	Lingwiesenstrasse 11/2	No. 3 Changi South Street 2

Lake Mary, FL 32746 	70825 Korntal – Münchingen	#01-01 Xilin Districentre Building B

USA	Germany	Singapore 486548

1.800.736.2771			+49.7150.97.97-400            	1.800.511.1360

support@faro.com	support@faroeurope.com	supportAP@faro.com

Disclaimer:  FARO Technologies, Inc. makes no warranty, express, implied or statutory, as to the material on this form, including any referenced third party.  The material contained on this form is for informational purposes only; such material is provided “AS IS” and with all faults, and your use of the material is at your own risk.